ST-270D Spin Dryer

Standard Operating Procedure Rev. 2026-01

Application

The ST-270D is designed for drying wafers without leaving water spots using cassettes. It has an automated program that will perform a quick DI water rinse and then dry the wafer(s) using heated nitrogen while spinning at high rpm. It is capable of handling 2 inch, 3 inch, 4inch and 6 inch wafers as well as pieces.

NRF Contact

David Hays

Table of Contents

1.	Safety information	3

2.	Prerequisites for use	3

3.	Sample Restrictions	4

4.	Start Up	4

5.	Tool Operation	4

6.	Ending The Session	6

Safety information

- Moving components – The tool has rotating parts at very high rpm’s. Never attempt to open the door while the rotor is still in motion.

Prerequisites for use

- User of the RSC facilities.

- Up-to-date hazardous waste training.

- Training by staff.

- User will be responsible for lost or damaged equipment due to negligence. Negligence means not following the SOP, not asking staff for assistance, intentionally damaging items, etc.

Sample Restrictions

The tool can only dry round substrates of 2”, 3”, 4” and 6”. Do not attempt to dry pieces or masks.

Use only the designated cassettes. Other cassettes may fit, but could cause the system to be unbalanced and lead to wafer breakage or damage to the tool.

Do not use this equipment as a rinser! It is a dryer only, and all substrates must be properly rinsed with DI water before loading. All loaded samples must be clean and free of chemicals.

Start Up

Log in to TUMI

Tool Operation

Login using TUMI system to activate tool and start billing period.

Open the door and place the proper rotor insert (2”, 3”, or 4”) for your substrate diameter in the slot.

There are labels stamped on the front of the inserts and the label must face the door.

When your wafers have finished rinsing, place the cassette in the opening of the insert with the H-bar to the back.

CAUTION        CAUTION

Only use the designated (marked) cassettes!

Close the door

Press the green, start button to begin the process

If you hear anything unusual; wafers breaking, unbalanced rotor, knocking, etc….press the stop button and contact staff.

When the process is finished, the rotor “should” go back to the upright position, the N2 will turn off and the door will unlock.

Open the door and remove the cassette. If rotor did not stop in upright position, simply rotate by hand and then remove the cassette.

Close the door

Place cassette back on top of the dryer

Log off the TUMI system

Ending The Session

Log off TUMI