# FEI Themis Z S/TEM

Electron microscope pixel array detector (EMPAD) Standard Operating Procedure Rev. 2026-01

Application

The FEI Themis Z S/TEM is a state-of-the-art Cs probe-corrected field emission gun scanning/transmission electron microscope capable of providing micro- to pico-scale structural, morphological, magnetic, and chemical information.  This document covers operation of the instrument for collecting 4D-STEM datasets using the electron microscope pixel array detector (EMPAD).  Please see separate SOP documents on the NRF instrument page regarding other applications including CS-corrected STEM (AKA “STEM”) and microprobe STEM (μP-STEM) mode operation.

NRF Contact

Nicholas G. Rudawski

# Table of Contents

1.	Safety	3

2.	Pre-requisites	3

3.	Voltage settings	3

4.	Scope of use	3

5.	Considerations for using STEM mode	4

6.	Considerations for using μP-STEM mode	4

7.	Setting up the STEM image	6

8.	Starting the EMPAD	7

9.	Background collection	10

10.	Centering CBED patterns on the EMPAD	11

11.	Search scan settings	13

12.	Virtual STEM detector settings	15

13.	Fine tuning (as needed)	17

14.	Acquiring a 4D-STEM dataset	18

15.	Finishing with the EMPAD	21

16.	Locating and transferring datasets	22

Safety

  - Due to stipulations of the instrument service contract as well as contamination concerns, analysis of radioactive and/or biological specimens with this instrument is strictly prohibited.

  - The EMPAD is radiation hard and very robust to high intensity electron irradiation so there is very little concern about damaging it during operation; that being said, it is still critical to stay at or below the threshold of 2 pA/pixel as this will avoid clipping the 32-bit dynamic range of the detector (discussed subsequently in greater detail).

Pre-requisites

  - Extensive familiarity with use of the Microscope Control user interface software and Velox imaging software for STEM imaging.

  - Proficient in independent operation of the FEI Themis Z S/TEM for imaging in STEM and/or μP-STEM modes.

  - NOTE: as per the prerequisites, some associated details and images that could otherwise be included here are omitted for purposes of clarity (“not shown; described elsewhere”).

Voltage settings

  - The EMPAD can be effectively used for any of the three voltage settings for the Themis; no modifications to the EMPAD are required when the microscope voltage is changed.

Scope of use

  - The EMPAD is used to collect so-called “4D-STEM” datasets where a 128×128 diffraction pattern (32-bit dynamic range) is collected at each pixel in a defined rectangular area scanned by a focused STEM probe; thus, each DP will be a convergent beam electron diffraction (CBED) pattern.  If the dataset is collected from the full (square) scanned area, it may have dimensions up to 512×512.

Considerations for using STEM mode

  - STEM mode is used to form sub 0.1 nm atomic-scale probes; thus, it is typically used for 4D-STEM datasets where atomic-resolution analysis is needed.

  - Due to the large values for the probe semi-angle of convergence (), there is little practical concern about exceeding the 2 pA/pixel EMPAD saturation threshold.

  - STEM mode can therefore be configured as it would be for normal atomic-resolution STEM imaging in preparation for 4D-STEM dataset collection with the EMPAD.

Considerations for using μP-STEM mode

  - μP-STEM is used to form nm-scale probes; thus, it is typically used for datasets where nm-scale (or larger) resolution analysis is needed.

  - Due to the small  values used in μP-STEM mode, there is greater concern about exceeding the 2 pA/pixel EMPAD saturation threshold.

  - The max probe current to avoid saturation (isat) can be calculated as a function of  in mrad (mrad) and the calibrated field of view in mrad (FOVmrad) for the CBED patterns collected by the EMPAD for the indicated camera length using the following formula:

isat= 2 pApixπ128 pix αmradFOVmrad2.

    - From this equation, it is evident that isat decreases rapidly with decreasing mrad and increasing FOVmrad (or decreasing camera length).

  - If it is not possible to set the probe current at or below isat for the specified conditions by modifying the spot number, the monochromator focus can be directly adjusted to obtain this (not shown; described elsewhere).

    - If there is a significant deviation from monochromator focus = –40, this can affect overall μP-STEM alignment.

    - Thus, it is best practice to first specify , then adjust the monochromator focus to modify the current (if needed) and finally perform μP-STEM alignment (not shown; described elsewhere).

  - Here is a table of the calibrated FOVmrad for the EMPAD DPs as a function of indicated camera length (L) for each voltage setting:

Setting up the STEM image

  - If needed, crystallographically align the specimen (not shown; described elsewhere).

  - Correct condenser astigmatism and optimize focus (not shown; described elsewhere).

    - Again, the goal being to keep the value of defocus as small as possible while keeping the image in focus.

  - If needed, adjust scan rotation (not shown; described elsewhere).

  - Adjust the magnification to obtain the desired field of view (FOV) for subsequent 4D-STEM dataset collection (not shown; described elsewhere).

    - NOTE: the FOV observed in Velox will typically be the same FOV over which a 4D-STEM dataset is collected using the EMPAD.

    - It will be possible to collect a 4D-STEM dataset from a subset of this area.

Starting the EMPAD

  - On the EMPAD PC desktop, select “PAD” to open the EMPAD GUI.

  - In the EMPAD GUI window, above “Project Settings”, select “New”.

    - In the “Create Project” window under “New Project Name”, enter a project name.

  - Above “Connections”, select “Go Online”.

  - Above “Detector”, select the box next to “Power on” to power up (and start cooling) the EMPAD.

  - Above “Temperature Control”, wait for the system to reach the T = –5 °C (negative 5 °C) setpoint before proceeding.

Background collection

  - Above “Detector”, verify “EMPAD inserted” is deselected; in Velox acquisition, blank the beam; no signal will reach the EMPAD.

  - Above “Exposure Control”, set “Exp. Time (ms)” = 1, Post Exp. Time (ms) = 0, and “Take Backgnd (N)” = 1000; select “Take Backgnd (N)” to acquire the background.

Centering CBED patterns on the EMPAD

  - In Velox acquisition, stop scanning and verify that the beam position marker is centered (not shown; described elsewhere).

  - In the EMPAD GUI above “Detector”, select “EMPAD inserted” to insert the EMPAD into the beam path.

  - In the EMPAD GUI under “GRAB_WINDOW”, select “Video On” to view the live CBED pattern.

  - The CBED pattern should now be evident under “GRAB WINDOW”.

    - Select “Cross” to mark the center of the window.

    - Adjust the “Intensity Scale” as appropriate to increase visibility (this has no effect on the CBED patterns in the final acquired 4D-STEM dataset).

    - In Microscope Control, use “Diffraction Alignment” to center the CBED pattern in the window (not shown; described elsewhere).

When finished, change the setting on the scan input control box from Velox (INT scan) to EMPAD (EDX scan).

NOTE: when viewed on the EMPAD the CBED patterns will be flipped upside down compared to when viewed on the FluCam.

Search scan settings

The “Search” scan setting is generally used for quickly visualizing the area to be scanned at low resolution.

  - Above “Scan Settings”, select “Search” to start scanning the area.

    - It is recommended to set “Scan size” = 1, which corresponds to the full square area.

    - It is recommended to set “Center” to X = 0.5 and Y = 0.5; this intuitively places the center position in the center of the area.

    - “Resolution” defines the resolution of the scanned area; both X and Y may be up to 512; e.g., if X = Y = 128, the scanned area will be the full square area with dimensions of 128×128; if X = 128 and Y = 64; the scanned area will be the full FOV in the X direction, but only half the FOV in the Y direction and centered vertically in the full square ROI.  X = Y = 128 is recommended for quickly scanning the full area (~16 s/frame)

Under “GRAB_WINDOW”, a reconstructed STEM image based on the “STEM1” virtual detector will be evident.

If any of the settings are changed, select “Focus Off” to stop scanning and then “Focus On” to restart scanning with the new settings.

Virtual STEM detector settings

NOTE: virtual detector settings have no impact on the CBED patterns in the final acquired 4D-STEM dataset; the reconstructed STEM image is simply used for quick visualization of the scanned area.

Under “GRAB_WINDOW”, verify a CBED pattern is not actively being acquired and scanning is stopped; “Video On” and “Focus On” should be indicated.

Select the “STEM1” virtual STEM detector and adjust the parameters (center, radii, etc.) as desired; when finished, select “Apply”.

Under “GRAB_WINDOW”, select “Focus On” to start acquiring a STEM image with the adjusted virtual detector parameters.

Fine tuning (as needed)

Due to the relatively long (minimum) dwell time of 1 ms per pixel used by the EMPAD, it tends to be difficult to accurately optimize focus/stigmation, scan rotation, magnification, and positioning of the specimen.

If any of these need to be adjusted, it is best to do this using the live HAADF image in Velox Acquisition.

    - If the EMPAD is inserted, under “Detectors”, deselect “EMPAD inserted” to retract it; otherwise, the EMPAD blocks the beam path to the HAADF detector.

    - Change the setting on the scan input control box from EMPAD (EDX scan) to Velox (INT scan).

    - Make all necessary adjustments using the live HAADF image in Velox (not shown; described elsewhere).

Once adjustments are complete, switch the scan input control box from Velox (INT scan) back to EMPAD (EDX scan) and reinsert the EMPAD into the beam path as described previously.

Acquiring a 4D-STEM dataset

Under “GRAB_WINDOW”, verify a CBED pattern is being actively viewed and scanning is currently stopped; “Video Off” and “Focus On” should be indicated.

Above “Scan Settings”, select “Acquire” and set the scan settings as described previously.

    - NOTE: due to the information limit associated with digitized data, there is essentially no effective benefit in acquiring a dataset with pixel size smaller than approximately ½ the probe size.

    - The pixel size of the dataset can be calculated by checking the measured FOV in Velox and dividing by the number of pixels in the X direction; e.g., for FOV ~ 1 μm and X = 512 pixels, x ~ 2 nm/pixel.

If desired, under “GRAB_WINDOW”, select “Focus On” to live STEM image (preview) with the specified scan settings (not shown); select “Focus Off” when finished.

  - Under “GRAB_WINDOW”, select “Take” to acquire the 4D-STEM dataset.

Under “GRAB_WINDOW”, a reconstructed STEM image based on the “Acquire” scan settings and “STEM1” virtual detector will be generated.

  - When the acquisition is complete, a dialogue box will appear asking for a name for the acquisition.

Finishing with the EMPAD

  - Above “Detector”, deselect “EMPAD inserted” to retract the EMPAD and “Power on” to power down the EMPAD (cooling will stop automatically).

  - Above “Connections”, select “Exit” to exit the EMPAD GUI.

Locating and transferring datasets

  - On the EMPAD PC taskbar, select “Files”; in the “Home” window, select “empad_projects”.

a screenshot of the EMPAD PC desktop

In the “empad_projects” window, select the project folder created when the EMPAD GUI was initially started.

The window will now have the project folder name; a folder containing the data for each acquired dataset (as named) will be shown.

Datasets may be transferred off the EMPAD PC using USB (preferably, 3.0) devices; keep in mind that 4D-STEM datasets are often very large with a full 512512 dataset being over 17 Gb in size.