## Slide 1: Photomask Creation

## Slide 2: Types of Photomasks used here at the RSC

## Slide 3
Nanofilm pre-coated photomasks.
Positive Photoresist – 5200A
Chrome oxide – 200A
Chrome – 800A
Soda lime glass – 2.3mm
Chrome Photomask
UV transmitting glass substrate with geometric patterns that are opaque to UV light. Used to transfer patterns to UV sensitive photoresist coatings.
Collimated UV Light Source
Chrome is wet etched and PR removed
Exposed and developed positive photoresist

## Slide 4
Positive Photoresist – 5200A
Iron Oxide 2000A
Soda lime glass – 2.3mm
Iron Oxide Photomask
Collimated UV Light Source
Iron Oxide is wet etched and PR removed
Exposed and developed positive photoresist

## Slide 5
Soda Lime Glass - Used for all RSC masks
4” and 5” available (2.3mm thickness)
Quartz not needed for most applications. Only when shorter wavelength exposure and lower thermal coefficient of expansion needed……
Photomask Blank type available at the RSC
Ref - http://www.mitani-micro.co.jp/en/mask/photo.html

## Slide 6
Low Reflective Chrome (chrome + chrome oxide)
          The Good
      Wet etched line edge roughness approx. 50nm
      1um minimum feature size
          The Bad
      Blocks alignment microscope light
Iron Oxide
  The Good
    Passes alignment microscope light
    Harder than chrome
    The Bad
    2um minimum feature size
    Wet etched line edge roughness approx. 300nm
    Harder to identify iron oxide side of the mask
Choosing Photomask Type

## Slide 7
RSC Heidelberg Exposure Tool
Heidelberg Laser Writer – 405nm laser write source

## Slide 8
  An ultrafast acousto-optic modulator turns the laser on/off at each pixel as the acousto-optic deflector sweeps the laser across the exposure area.
  The stage scans in the Y axis during exposure.
  The stage position resolution is 10nm.
Basic Optics configuration
Heidelberg Direct Write Laser System
Y Axis
X Axis

## Slide 9
Heidelberg Direct Write Laser System
Can expose flat substrate
Repeatable and accurate
Very good focus control via autofocus system
      20mm head for ≥ 5um features - DOF=60um
      4mm head for ≥ 1um features – DOF=1.6um
      Very good overlay/alignment (100nm for 4mm mode)
Design changes on the fly

## Slide 10
Consider specs for Heidelberg exposed photomasks
20mm Mode (refers to the focal distance from lens)
Minimum feature size = 5.0 microns
edge roughness = 180nm-3σ,
CD uniformity = 440nm-3σ
Pixel Size = 1000nm
Typical Time = 1 hour , 80mm² area
4mm Standard Mode
Minimum feature size = 1.0 microns
edge roughness = 70nm-3σ,
CD uniformity = 100nm-3σ
Pixel Size = 400nm
Typical Time = 8 hours, 80mm² area
4mm HQ Mode
Minimum feature size = 1.0 microns
edge roughness = 30nm-3σ,
CD uniformity = 50nm-3σ
Pixel Size = 200nm
Typical Time = 16 hours, 80mm² area

## Slide 11
CAD Software for Pattern Creation

## Slide 12
CAD file formats
GDSII strongly encouraged - The RSC provides use of “Layout Editor” software. Shareware version also available
  Klayout is another option – open source
Autocad
  may also be used but not recommended. If you insist…..rules must be followed.
    Save all files as “AutoCAD revision R12/LT2 DXF format.
    .dwg files are NOT supported.
    Do not open a .dxf file in a GDSII software such as “Layout Editor” or Ledit” and save it as a GDSII file.
    Set units to microns
    Use only the following entities: CIRCLE, POLYLINE / LWPOLYLINE and TEXT. LWPOLYLINE is preferred since it’s 2D.
    Send the file to Bill Lewis to be converted to GSDII with Linkcad 7 software.
    Read Section 5.2 of “NRF Heidelberg DWL66fs SOP” for entire list of rules……

## Slide 13
What typically happens when Autocad or Solidworks is used?
90% of all .dxf files I receive have problems
    line segments
    Nested Polygons
    Wrong scale….meters, millimeters, nanometers etc..
    Large unusable file sizes
    Intersecting polygons
    Text labels instead of polygons
    Remedy for 50% of issues - convert to GDSii using Linkcad 7 software

## Slide 14
General Rules for ALL designs
No intersecting polygons
1
2
3
4
1
2
3
4
5
6
1
2
3
1
2
3
NO
YES
YES
No line segments
No nesting polygons
1
2
3
1
4
2
1
2
2
1
1
2
1
2
NO
NO
YES

## Slide 15
Heidelberg software inverts pattern
Mirror in X
Easy to create
hard to create
Text labels are not polygons
Mask polarity should not be a consideration when create your drawing
Mask aligner masks require X mirroring if topside aligned (i.e. looking through the glass side of the mask)
General Rules for ALL designs
Align die in straight grids if you will be using the dicing saw and leave room in street for dicing blade i.e. >200um.
Overlap is OK

## Slide 16
Entire device in one gsdii
use “layers” to draw each process layer of the device
“cell” function and hierarchy
Only one circle, rectangle and triangle.
“Layout Editor” GDSii CAD Software
Cell=circle (layer 1)
Cell=rectangle (layer 2
Cell=triangle (layer 3)
Cell=circle array
Cell=rectangle array
Cell=triangle array
Cell=Photomasks
The rest are all “cell references”.
Sub-sub cell changes impact all downstream cells

## Slide 17
Freeware version will only save to 1000 polygons
Mouse operations
Coordinate display
Snap to Grid
Draw functions (box, free polygon)
Snap to point
Circle – limiting vector points
Selecting polygons and points
“properties” settings
Layer controls for mask device layers
Layer colors/hatching
Boolean operation
“Layout Editor” Software Demo

## Slide 18
Inserting single cell reference
Inserting Array cell reference
Flatten cell to allow edit of array
Moving polygons or cell refs. Use “move by” function = “right click, move, m, shift+left click”
Example of nested polygons
Use separate non-working layer to make layout boundaries
Software Demo

## Slide 19
Design Layout Affects Cost

## Slide 20
Requesting Masks
or
Training to make your own mask

## Slide 21
Making your own masks
  Sign up for Heidelberg training
  Create your design using Layout Editor
  Submit the design via email to walewis@ufl.edu
  Request masks from NRF Staff, email walewis@ufl.edu the day before.
  Expose, develop, etch, inspect, strip PR.
Staff Making your mask
  Create your design using Layout Editor
  Submit the design via email to walewis@ufl.edu for review
  Fill out “Staff Processed Photomask Request” form which is located on the “Docs” screen of the RSC Heidleberg web page. Email form to walewis@ufl.edu
  If everything looks good, “Request Service” on the RSC web page.

## Slide 22
Information Resources

## Slide 23
Mask Cost Estimator

## Slide 25
When the design is inverted during conversion – use “Frame” parameter.

## Slide 26
Do you really need a Photomask?

## Slide 27
Start
Will layer be reproduced > 3 times
Min Feature sizes <1um?
Yes
Min Feature size 1-2um
Raith E-beam writer
Expose on Mask Aligner
Make mask using the Heidelberg Laser Writer
Yes
No
Yes
No
Min Feature size 2-5um
Consult with NRF Staff
No
Yes
Heidelberg Direct Write Laser Writer
No
Mask or Direct Write (<2um thickness and 405nm sensitive materials ONLY)

## Slide 28
Photomask Design Considerations

## Slide 29
Feature resolution
  minimum resolution? It depends
  Suss MA6 Mask Aligner = 1um with good planning
  whole 4” flat substrates are your best bet
  Process bias considerations
  Consult with NRF Staff if you are planning on printing <4um features.
Overlay alignment tolerances
  Consider layer-to-layer alignment tolerance. You should design your pattern with a tolerance of at least 2um overlay error when possible.
Mask Aligner Expectations

## Slide 30
5” Mask
Alignment Microscope Objective Lens
Area of Microscope Travel
Typical Aligner Overlay Setup
Range of motion for the alignment microscope is limited, so
alignment marks must be within
A specific window on the mask.
Allowable Alignment Mark Positions
The closest spacing for the Suss MA6 5x and 20x objectives is 40mm. The MA6 is also equipped with a 10x “offset” objective. The min spacing for it is 8mm.
EV620 Objective range
•Top side objective travel range: x direction 30 - 150mm separation, y direction +-75mm; z direction +-5mm
•Bottom side objective travel range: x direction 30 - 100mm separation (8-100mm optional); y direction +-12mm, z direction +-5mm
100 mm
70 mm
40 mm

## Slide 31
Alignment Marks
Vernier arrangements for alignment marks are preferred. Staff can provide a GDSII file for the structure below.
The step pitch difference of the fingers are used to measure mis-alignment.
Layer 1
Layer 2

## Slide 32
Alignment Marks – (cont)
Chrome mask alignment marks
  You will have to look through the mask via the alignment microscopes to hunt for the alignment marks on your device. The field of view of the CCD image for the 5X objective is 850um. It’s best to have as much of that field of view clear as possible so you can find the wafer alignment marks.
 | Alignment Objective | Approximate Split-field CCD window in um
Suss MA6 | 5X | 850
 | 10X | 380
 | 20X | 210
EVG 620 | top 10X | 350
 | bot 10X | 350

## Slide 33
Will you be able to see alignment marks on your device?
Consider the affect of processing steps between mask layers. How will the addition of films affect the visibility in the alignment microscope?
Top down light diffracts off the step edge to create edge shadow
Thin conformal films are usually visible using top by top alignment down microscope
Thin Aluminum
Thick Aluminum
Alignment Mark may not be visable

## Slide 34
Process considerations for masks

## Slide 35
Mask Layout for Plasma Etch tools
Safe area for pattern 90mm standard. 75mm for DRIE
No pattern zone – edge exclusion width for STS DRIE= 12.5mm------needs PR protection
Edge exposure or PR removal for Unaxis etcher 8mm around edge----No PR. Design this into your mask!
Etcher Electrode
wafer
Unaxis ICP Wafer Clamp

## Slide 36
Feature Size Bias
  Litho Bias
    Mask Size vs. Actual PR Size
  Plasma Etch Bias
    Plasma etch erosion of photoresist mask sidewalls may affect slope and roughness of the target material sidewall.
    It may change final feature size.
    The mask may need to be “biased” to compensate for the feature size change.
Mask Bias

## Slide 37
  Wet Etch Undercut
1um etch = 1um undercut
Resist
  Lateral loss is almost as much as vertical.
  Mask writers usually have an option to bias the design size to compensate. On the Heidelberg Laser Writer, “spotsize correction” X and Y can be used to do this within limits as determined by pixel size.
How can I determine bias for my process?
Mask Bias