BSC-100 Spinner

Standard Operating Procedure Rev. 2026-01

Application

The BSC-100 spinner is used to manually apply difficult to clean materials such as SU8, spin on glass, ITO, etc. The spin speed, time and ramp can be programmed up to five steps. It can be used for small pieces to wafers up to 100mm and has a maximum spin speed of 6000rpm.

NRF Contact

David Hays

Table of Contents

1.	Safety information	3

2.	Prerequisites for use	3

3.	Operation	4

3.1.	Start Up	4

3.2.	Chuck Selection	4

3.3.	Operating modes	5

3.4.	Cleaning	6

4.	Ending The Session	6

Safety information

- Moving Components – The User should be aware at all times of the moving components associated with this tool. The spinner lid must be closed while processing your sample.

Prerequisites for use

- User of the RSC facilities.

- Up-to-date hazardous waste training.

- Training by staff.

- User will be responsible for lost or damaged equipment due to negligence. Negligence means not following the SOP, not asking staff for assistance, intentionally damaging items, etc.

Operation

Start Up

Log in to TUMI

Chuck Selection

The proper chuck must be used or your sample may be broken, or cause damage to the motor. Your sample MUST be larger than the o-ring or photoresist will be sucked into the motor and damage it.

For fragments, you will need to use the chuck with the smallest o-ring.

Two inch wafers must be used with the chuck with the larger o-ring. DO NOT USE THE LARGE, FLAT CHUCK (55mm) because the wafer is smaller than the vacuum area and resist will be sucked into the motor.

4 inch wafers

- Use the large, flat chuck shown below

Operating modes

- There are two modes of operation…Easy and Normal. When you first log in, the tool will be in Easy mode which means it is a single spin step only, and will display “Easy Mode” on the screen.

Single step operation (Easy mode)

- Remain in “Easy Mode”

- Press and hold the ‘Set’ button for 3s to enter edit mode

- Use the ‘Left/Right’ arrows to move the cursor and the ‘Up/Down’ arrows to change the values

- Press and hold the ‘Set’ button for 3s to exit edit mode

- Press the ‘Pump’ button to turn on vacuum and ensure that it is greater than 20 in. Hg, otherwise your sample may fly off.

- Press the start button  to start the process. This button can also be used to stop the process before the time ends, if needed.

Multi-step operation (Normal mode)

- Press the ‘Mode’ button to switch to standard mode

- Press the ‘Left/Right’ arrows to cycle through the 5 recipes, labeled A through E.

- Press and hold the ‘Set’ button for 3s to enter edit mode

- Use the ‘Left/Right’ arrows to move the cursor and the ‘Up/Down’ arrows to change the values for that step

- Press the ‘Set’ button to move to the next step and repeat step (4) to edit

- Press and hold the ‘Set’ button for 3s to exit edit mode

- Press the ‘Pump’ button to turn on vacuum and ensure that it is greater than 20 in. Hg, otherwise your sample may fly off.

- Press the start button  to start the process. This button can also be used to stop the process before the time ends, if needed.

Cleaning

Grab the bowl and lift, to remove it from the spinner. Do not try to clean or spray solvents in the bowl with it still on the spinner!

Use isopropanol and wipes to remove the bulk of the material. If the material is dried in the bowl (i.e. it’s been more than 1 hour since it was dispensed) use AZ EBR 7030 to clean the bowl.

Wipe the entire bowl and lid with clean wipes.  Apply more solvent and wipe again until the both are extremely clean. The final clean should be with acetone. There will be no more color changes to the wipe when the bowl is clean.

SU8 USERS…after removing the bulk of SU8 with propanol and acetone, your final cleans should be with SU8 Thinner.

Wipe down the vacuum chuck with isopropanol and then acetone.

Place contaminated wipes in the solvent waste can.

FAILURE to clean the bowl properly will cause you to lose privileges on this spinner.

Ending The Session

Log off TUMI